## GUM steps
The Presentation of the results of Measurements can be divided into eight steps
1. Design a measurement equation (model of evaluation) for determination of the measurand including all input quantities that significantly contribute to the measurement result (value and associated MU).
The output quantity is stated and the model of evaluation developed as an algebraic equation combining all the input quantities to calculate the value of the output quantity.
2. Determine the estimated value for all input quantities
3. Determine for all input quantities the associated standard uncertainty using Type A evaluation (statistical method) or Type B evaluation (any other method).
4. Determine the covariances for all input quantities with correlated MUs.
5. Calculate measurement results from input quantities with the model of evaluation.
6. Determine the combined standard MU from the standard MUs and covariances of the input quantities and from the related sensitivity coefficients.
7. Determine the effective DOF and the related value for the coverage factor for the intended confidence level.
8. Report the result of the measurement together with the associated expanded MU

## Sphere Inc-Inc
Calibration of Luminous Flux by and Integrating Sphere
Incandescent reference / Incandescent test lamp (secondary standard)
current set using a voltmeter to measure the current through a calibrated shunt, R
broadband detector with lumen photpic filter used to obtain signals
Model
variables relating to test lamp have no subscript
subscript 'R' refers to reference lamp standard
subscript 's' refers to sphere
Table
No | Quantity Xi | Symbol | Unit | Value xi | Abs. std. uncertainty u(xi) | Type | DOF n | Abs. sensitivity ci | Abs. contribution ui(y) | Relative contribution urel(y)
3 | sphere spatial response factor -test lamp | DSf |  | 0.001 | 0.0012 | B | ¥ | -1 | -0.0012 | -0.0012012
4 | relative aging coeff.-test lamp | g | 1/h | 0.0003 | 0.0003 | B | ¥ | -0.1 | -3e-05 | -3.0031e-05
5 | burning time-test lamp | Dt | h | 0.1 | 0 | B | ¥
6 | correction factor -test lamp | corS |  | 0.99897 |  |  | ¥ |  | 0.0012004 | 0.0012016
9 | sphere spatial response factor-REF | DSfR |  | 0 | 0.0005 | B | ¥ | -1 | -0.0005 | -0.000501
14 | relative aging coeff.-REF | gR | 1/h | 0.0002 | 0.0001 | B | ¥ | -10 | -0.001 | -0.001002
15 | burning time-REF | DtR | h | 10 | 0 | B | ¥
12 | correction factor-REF | corSR |  | 0.998 |  |  | ¥ |  | 0.001118 | 0.0011203
13 | correction factor-REF | corSR |  | 0.998 | 0.001118 | A | ¥ | 7026.387 | 7.8557 | 0.0011203
14 | correction factor-test lamp | corS |  | 0.99897 | 0.0012004 | A | ¥ | -7019.5745 | -8.4261 | -0.0012016
15 | luminous flux-REF | FR | lm | 2800 | 4.8 | B | ¥ | 2.5044 | 12.0211 | 0.0017143
16 | photometer-source | y | V | 4.579 | 0.0085 | A | 9 | 1531.4117 | 13.017 | 0.0018563
17 | photometer-REF | yR | V | 1.765 | 0.0007 | A | 9 | -3972.9945 | -2.7811 | -0.0003966
18 | photometer AUX-test lamp | yA | V | 3.1234 | 0.001 | A | 9 | -2245.0967 | -2.2451 | -0.00032016
19 | photometer AUX-REF | yAR | V | 3.1243 | 0.001 | A | 9 | 2244.4497 | 2.2444 | 0.00032007
20 | shunt voltage-test lamp | U | V | 0.88 | 0.00012 | A | 9 | -55461.201 | -6.6553 | -0.00094909
21 | shunt voltage-REF | UR | V | 0.58 | 7.5e-05 | A | 9 | 86082.455 | 6.4562 | 0.00092069
22 | current-lumen exponent-test lamp | mJ |  | 6.96 | 0.38 | B | ¥ | -3.2844 | -1.2481 | -0.00017798
23 | current-lumen exponent-REF | mJR |  | 7.12 | 0.22 | B | ¥ | -30.8616 | -6.7896 | -0.00096823
26 | spetral mismatch-test lamp | F |  | 1 | 0.001 | B | ¥ | -7012.3412 | -7.0123 | -0.001
27 | spectral mismatch-REF | FR |  | 1 | 0.0005 | B | ¥ | 7012.3342 | 3.5062 | 0.0005
28 | shunt resistor | R | Ohm | 0.10004 | 6e-06 | B | ¥ | -499068.6089 | -2.9944 | -0.00042702
29 | cal. factor current DVM | c |  | 1.0003 | 4.1e-05 | B | ¥ | 559.9575 | 0.022958 | 3.274e-06
32 | nominal current-source | J | A | 8.795 | 0
33 | nominal current-REF | JR | A | 5.825 | 0 |  |  |  |  | 0.0026315
34 | luminous flux - test lamp | F1 | lm | 7012.3342 |  |  | 123.691 |  | 25.8561 | 0.0036872
[6 empty] | found from approximation formula: k = | 2.03 |  | 52.488 | 0.0074851
[7 empty] | variance - Test Lamp |  |  | 0.2631%
[7 empty] | variance - REF |  |  | 0.2583%
𝑐𝑜𝑟𝑆=(1−𝑠𝑓−𝑡)
〖𝑐𝑜𝑟𝑆〗_𝑅=(1−〖∆𝑠𝑓〗_𝑅−𝛾_𝑅 〖∆𝑡〗_𝑅 )

=_𝑅·𝑦/𝑦_𝑅 ·𝑦_𝐴𝑅/𝑦_𝐴 ·((𝑈·𝑐)/(𝐽·𝑅))−𝑚_𝐽·((𝑈_𝑅·𝑐)/(𝐽_𝑅·𝑅)) 𝑚_𝐽𝑅·(𝐹_𝑅/𝐹)·〖𝑐𝑜𝑟𝑆〗_𝑅/𝑐𝑜𝑟𝑆

## Sphere Inc-SSL
Measurement of Luminous Flux by and Integrating Sphere
Incandescent reference / SSL or CFL test lamp
spectral detection
current read as DVM voltage with a shunt resistor, R
Model
test lamp variables do not have subscripts
subscript 'R' refers to reference lamp standard
subscript 's' refers to sphere
Table |  |  |  | manual entries are highlighted in blue
No | Quantity Xi | Symbol | Unit | Value xi | Abs. std. uncertainty u(xi) | Type | DOF n | Abs. sensitivity ci | Abs. contribution ui(y) | Relative contribution urel(y)
1 | ambient temp. difference | DTa | °C | 2 | 1 | A | 29 | 0.002 | 0.002 | 0.0019921
2 | ambient temp. coeff. | a | 1/K | 0.002 | 0.001 | A | 29 | 2 | 0.002 | 0.0019921
3 | sphere spatial response factor | Dsf |  | 0 | 0.0012 | B | ¥ | -1 | -0.0012 | -0.0011953
4 | stray light | S(q) |  | 0 | 0.005 | B | ¥ | -1 | -0.005 | -0.0049802
5 | relative aging coeff. | g | 1/h | 0.0003 | 0.0003 | B | ¥ | -0.1 | -3e-05 | -2.9881e-05
6 | burning time | Dt | h | 0.1 | 0 | B | ¥
7 | correction factor -test lamp | corS |  | 1.004 |  |  | 2150 |  | 0.0058686 | 0.0058454
8 | ambient temp. difference-REF | DTaR | °C | 2 | 1 | A | 29 | 0.00015 | 0.00015 | 0.00015002
9 | ambient temp. coeff. -REF | aR | 1/K | 0.00015 | 2e-05 | B | ¥ | 2 | 4e-05 | 4.0004e-05
10 | sphere spatial response factor-REF | DsfR |  | 0 | 0.0005 | B | ¥ | -1 | -0.0005 | -0.00050005
11 | stray light-REF | SR(q) |  | 0 | 0.002 | B | ¥ | -1 | -0.002 | -0.0020002
12 | relative aging coeff.-REF | gR | 1/h | 0.0002 | 0.0001 | B | ¥ | -2 | -0.0002 | -0.00020002
13 | burning time-REF | DtR | h | 2 | 0 | B | ¥
14 | correction factor-REF | corSR |  | 0.9999 |  |  | 1066138 |  | 0.002077 | 0.0020772
15 | correction factor-REF | corSR |  | 0.9999 | 0.002077 | A | 1066138 | 909.4407 | 1.8889 | 0.0020772
16 | correction factor-test lamp | corS |  | 1.004 | 0.0058686 | A | 2150 | -905.7849 | -5.3157 | -0.0058456
17 | luminous flux-REF | FR | lm | 1715 | 8.6 | B | ¥ | 0.53023 | 4.56 | 0.0050146
18 | photometer-test lamp | y | lm | 920 | 9.6 | A | 9 | 0.98842 | 9.4889 | 0.010435
19 | photometer-REF | yR | lm | 1715 | 5 | A | 9 | -0.53024 | -2.6512 | -0.0029155
20 | photometer - AUX test lamp | yA | lm | 790 | 2 | A | 9 | -1.1511 | -2.3022 | -0.0025317
21 | photometer-AUX REF | yAR | lm | 805 | 1 | A | 9 | 1.1296 | 1.1296 | 0.0012422
22 | measured voltage-test lamp | U | V | 119.99 | 0.012 | A | 29 | 0.0078927 | 9.4713e-05 | 1.0415e-07
23 | cal. factor AC voltage | cU |  | 0.9986 | 0.00104 | B | ¥ | 910.6246 | 0.94705 | 0.0010415
24 | voltage-lumen expoent - test lamp | mU |  | 0.5 | 0.1 | A | 9
25 | shunt voltage-REF | UR | V | 0.0849 | 6e-06 | A | 29 | 76261.1356 | 0.45757 | 0.00050318
26 | cal. factor DC voltage | cR |  | 0.9965 | 4.3e-05 | B | ¥ | 6497.3109 | 0.27938 | 0.00030724
27 | shunt resistor | R | Ohm | 0.099999 | 7.1e-05 | B | ¥ | -64746.6232 | -4.597 | -0.0050553
28 | current-lumen exponent-REF | mJR |  | 7.12 | 0.22 | B | ¥ | -3.181 | -0.69983 | -0.00076959
29 | nominal voltage - source | U0 | V | 120 | 0
30 | nominal current-REF | JR | A | 0.849 | 0 |  |  |  |  | 0.012226
29 | luminous flux- test lamp | F | lm | 909.3498 |  |  | 35 |  | 13.3817 | 0.014716
[6 empty] | found from approximation formula: k = | 2.09 |  | 27.9677 | 0.030756
[8 empty] | variance - Test Lamp |  | 1.223%
[8 empty] | variance - REF |  | 0.8191%
=_𝑅·𝑦/𝑦_𝑅 ·𝑦_𝐴𝑅/𝑦_𝐴 ·((𝑈·𝑐_𝑈)/𝑈_0 ) 𝑚_𝑈·((𝑈_𝑅·𝑐_𝑅)/(𝑅∙𝐽_𝑅 )) 𝑚_𝐽𝑅·〖𝑐𝑜𝑟𝑆〗_𝑅/𝑐𝑜𝑟𝑆
corS
corS=(1+〖∆𝑇〗_𝛼−𝑠𝑓−𝑡)
𝑐𝑜𝑟𝑆=(1+〖∆𝑇〗_𝛼−𝑠𝑓−𝑆(𝜃)−𝑡)
〖𝑐𝑜𝑟𝑆〗_𝑅=(1+𝛼_𝑅 〖∆𝑇〗_𝛼𝑅−〖∆𝑠𝑓〗_𝑅−𝑆_𝑅 (𝜃)−𝛾_𝑅 〖∆𝑡〗_𝑅 )

## Gonio Intensity
Measurement of Luminous Intensity by and a Goniometer
Incandescent reference / Any type of test lamp
broadband detector with lumen filter
current read by DVM using calibrated shunt
Model
test lamp variable do not have subscripts
subscript 'R' refers to reference lamp standard
Table |  |  |  | manual entries are highlighted in blue
No | Quantity Xi | Symbol | Unit | Value xi | Abs. std. uncertainty u(xi) | Type | DOF n | Abs. sensitivity ci | Abs. contribution ui(y) | Relative contribution urel(y)
1 | amb.temp. difference - test lamp | DTa | °C | 2 | 1 | A | 29 | 0.0007 | 0.0007 | 0.00069974
2 | rel.temp.coeff. - test lamp | a | 1/K | 0.0007 | 0.0003 | B | ¥ | 2 | 0.0006 | 0.00059978
3 | relative dist. variation - test lamp | Dd/d |  | 0 | 0.0003 | B | ¥ | 2 | 0.0006 | 0.00059978
4 | angular tilt - test lamp | h(q) |  | 0 | 0.0015 | B | ¥ | 1 | 0.0015 | 0.0014994
5 | angular turn - test lamp | k(j) |  | 0 | 0.00065 | B | ¥ | 1 | 0.00065 | 0.00064976
7 | stray light - test lamp | S(q) |  | 0.001 | 0.001 | B | ¥ | -1 | -0.001 | -0.00099963
6 | rel.aging coeff. - test lamp | g | 1/h | 0.0003 | 0.0003 | B | ¥ | -0.1 | -3e-05 | -2.9989e-05
8 | burning time - test lamp | Dt | h | 0.1 | 0 | B | ¥
9 | correction factor - test lamp | corS |  | 1.0004 |  |  | 2880 |  | 0.0022098 | 0.002209
10 | amb.temp. difference - REF | DTaR | °C | 2 | 1 | A | 29 | 0.00015 | 0.00015 | 3.2758e-05
11 | rel.temp.coeff. - REF | aR | 1/K | 0.00015 | 2e-05 | B | ¥ | 2 | 4e-05 | 1.4286e-07
12 | relative distance variation - REF | DdR/dR |  | 0 | 0.0005 | B | ¥ | 2 | 0.001 | 0.0010057
13 | angular tilt - REF | hR(q) |  | 0 | 0.0002 | B | ¥ | 1 | 0.0002 | 0.00022727
14 | angular turn - REF | kR(j) |  | 0 | 0.0004 | B | ¥ | 1 | 0.0004 | 0.00039985
16 | stray light - REF | SR(q) |  | 0.005 | 0.001 | B | ¥ | -1 | -0.001 | -3.5714e-06
15 | rel.aging coeff. - REF | gR | 1/h | 0.0002 | 0.0001 | B | ¥ | -5 | -0.0005 | -0.00050287
17 | burning time - REF | DtR | h | 5 | 0 | B | ¥
18 | correction factor source - REF | corSR |  | 0.9943 |  |  | 350645 |  | 0.0015729 | 0.000226
19 | correction factor - REF | corSR |  | 0.9943 | 0.0015729 | A | 350645 | 672.5338 | 1.0578 | 0.0015819
20 | correction factor -test lamp | corS |  | 1.0004 | 0.0022098 | A | 2880 | -668.4563 | -1.4772 | -0.002209
21 | luminous intensity - REF | IR | cd | 280 | 0.68 | B | ¥ | 2.3882 | 1.624 | 0.0024286
22 | photometer signal - test lamp | y | V | 4.579 | 0.0085 | A | 9 | 146.0363 | 1.2413 | 0.0018563
23 | photometer signal - REF | yR | V | 1.765 | 0.0007 | A | 9 | -378.8671 | -0.26521 | -0.0003966
24 | shunt voltage | U | V | 0.88 | 0.00012 | A | 9 | -5288.8132 | -0.63466 | -0.00094909
25 | shunt voltage - REF | UR | V | 0.58 | 7.5e-05 | A | 9 | 8208.8742 | 0.61567 | 0.00092069
26 | current-intensity exponent | mJ |  | 6.96 | 0.38 | B | ¥ | -0.31321 | -0.11902 | -0.00017798
27 | current-intensity exponent - REF | mJR |  | 7.12 | 0.22 | B | ¥ | -2.943 | -0.64746 | -0.00096823
28 | Spectral mismatch - test lamp | F | K | 2900 | 25 | B | ¥ | -0.2306 | -5.7651 | -0.0086213
29 | Spectral mismatch - REF | FR | K | 2800 | 10 | B | ¥ | 0.23882 | 2.3882 | 0.0035714
30 | detector non-linearity-test lamp | H(q) |  | 1.002 | 0.002 | A | 9 | -667.3683 | -1.3347 | -0.001996
31 | detector non-linearity-REF | HR(q) |  | 0.995 | 0.002 | A | 9 | 672.0607 | 1.3441 | 0.0020101
32 | Distribution sample - test lamp | Ds |  | 1 | 0.01 | B | ¥ | -668.7672 | -6.6877 | -0.010001
33 | Distribution sample - REF | DsR |  | 1 | 0.001 | B | ¥ | 668.7004 | 0.6687 | 0.001
34 | shunt resistor | R | Ohm | 0.10004 | 6e-06 | B | ¥ | -1069.4714 | -0.0064168 | -9.596e-06
35 | cal. factor - DVM | c |  | 1.0003 | 4.1e-05 | B | ¥ | 106.9578 | 0.0043853 | 6.5579e-06
36 | nominal current of test lamp | J | A | 8.795 | 0
37 | nominal current - REF | JR | A | 5.825 | 0 |  |  |  |  | 0.01355
38 | intensity at nominal current | I | cd | 668.7004 |  |  | 31018 |  | 9.8217 | 0.014688
[6 empty] | found from approximation formula: k = | 2 |  | 19.6433 | 0.029375
[8 empty] | variance | test lamp | 1.355%
[8 empty] | variance | REF | 0.5668%
𝑐𝑜𝑟𝑆=(1+〖〖∆𝑇〗_𝛼+2𝑑〗∕𝑑+ℎ(𝜃)+𝑘()−𝑆()−𝑡)
〖𝑐𝑜𝑟𝑆〗_𝑅=(1+〖𝛼_𝑅 〖∆𝑇〗_𝛼𝑅+〖2∆𝑑〗_𝑅〗∕𝑑_𝑅 +ℎ_𝑅 (𝜃)+𝑘_𝑅 ()−𝑆_𝑅−𝛾_𝑅 〖∆𝑡〗_𝑅 )
𝐼=𝐼_𝑅·𝑦/𝑦_𝑅 ·((𝑈·𝑐)/(𝐽·𝑅))−𝑚_𝐽·((𝑈_𝑅·𝑐)/(𝐽_𝑅·𝑅)) 𝑚_𝐽𝑅·𝐹_𝑅/𝐹·〖𝐷𝑠〗_𝑅/𝐷𝑠(𝐻_𝑅 (𝜃))/(𝐻(𝜃))∙〖𝑐𝑜𝑟𝑆〗_𝑅/𝑐𝑜𝑟𝑆